## Slide 1: Tutorial: Film Dosimetry Analysis
Csaba Pinter
Laboratory for Percutaneous Surgery, Queen’s University, Kingston, ON, Canada

## Slide 2: Learning objective
This tutorial demonstrates how to perform a film dosimetry analysis with the streamlined “slicelet” in 3D Slicer.

## Slide 3: Overview
Install required software
Download test data
Perform calibration
Evaluate calibration

## Slide 4: 1/1. Install 3D Slicer
Download 3D Slicer preview from https://download.slicer.org/
Find the downloaded 3D Slicer package on your computer
Follow the usual steps to install an application
  Different for each operating system

## Slide 5: 1/2. Install SlicerRT extension
or

## Slide 6: 1/3. Install SlicerRT extension

## Slide 7: 1/4. Install FilmDosimetryAnalysis extension

## Slide 8: 2/1. Download test data
Open in browser: https://pocus.cs.queensu.ca/#collection/5cc8811601d4930406c40465/folder/5d4b1bae01d4930406c46b2c

## Slide 9: 2/2. Where did this data come from?
Calibration films were irradiated in well characterized radiation fields (10x10 cm2 at dmax) to a range of doses

## Slide 10: 2/3. Where did this data come from?
Treatment planning using Eclipse:
  Dual arc delivery FSRT VMAT (4 Gy delivered to mets, plus 2 Gy SIB)
  Delivery to a cylindrical phantom with film at isocenter plane
Outputs: Planning dose volume, Structures

## Slide 11: 2/4. Where did this data come from?
Irradiation day:
  Quarter sheets of Gafchromic EBT3 film used for calibration points
  Full sheet irradiation for experimental/measured film (FSRT VMAT irradiation)

## Slide 12: 2/5. Where did this data come from?
CCSEO Film Imager
0 cGy
50 cGy
100 cGy
200 cGy
300 cGy
400 cGy
500 cGy
600 cGy
Empty Light Field
FSRT VMAT Delivery
Calibration Set
Film imaged 24 hrs post-irr.

## Slide 13: 3/1. Open Film Dosimetry slicelet
Once in the ‘Film Dosimetry Analysis’ module, click ‘Show slicelet’

## Slide 14: 3/2. Load film images

## Slide 15: 3/3. Assign dose levels to films

## Slide 16: 3/4. Define calibration region
Click Add region, then click in the center of the film then to the region corner to define a region of interest (ROI).

## Slide 17: 3/5. Run calibration

## Slide 18: 3/6. Calibration ready
Calibration function has been determined, of the form:
Dose = a + bOD + cODn

## Slide 19: 3/7. Load experimental DICOM data
Select Film Dosimetry Dataset folder.
Then select the DICOM folder and click ‘Import’.
Then, when prompted click Add link.

## Slide 20: 3/8. Load experimental DICOM data

## Slide 21: 3/8. Load experimental film
Select Film Dosimetry Dataset folder that was just downloaded.
Then select the Images folder. Select the experimental.png file and click ‘Open’

## Slide 22: 3/9. Set experiment properties
Film resolution (mm/px) = 0.423
Define film plane and axis to be exported from DICOM dose distribution
Assign data as shown

## Slide 23: 3/10. Apply calibration
Apply calibration to convert film from optical density to dose
Note: Calibration function can be imported from a previously saved file, or can be input manually in this step.

## Slide 24: 4/1. Register film to plan
1. Pre-align film with the rotation and flip buttons
2. Click Perform registration to perform an intensity based registration between film and Eclipse dose plane
Grayscale: Calibrated film measurement
Greenscale: Eclipse dose plane

## Slide 25: 4/2. Dose comparison
Dose maps are automatically assigned
Select 74: RTSTRUCT Structure set. Choose ‘Brain’ structure.
Set dose value to 400 cGy (prescription)
Create new gamma volume
Set dose threshold to 0%

## Slide 26: 4/3. Gamma comparison done

## Slide 27: 4/4. Line profile

## Slide 28
Congratulations!