## Slide 1: Example parameter sets for automatic registration for different scenarios in 3D Slicer
Scenarios:
CVH phantom data for Varian (2012-09-10)
  Success rate on sample images: 100% (no visible misalignment after registration), fully automatic
Odette Cancer Centre data

## Slide 2: CVH phantom data for Varian: Planning MRI to guidance MRI
Generic registration (BRAINS)
  Rigid (6 DOF)
  Number of samples=10000
  Minimum step length=0.05
  Transform scale=100000.0
  Relaxation factor=0.9
  Maximum step length=10.0
Results
  Success rate on sample images: 100% (no visible misalignment after registration), fully automatic
  Computation time: under 5 seconds

## Slide 3: CVH phantom data for Varian: CT to CBCT
Generic registration (BRAINS)
  Rigid (6 DOF)
  Number of samples=100000
  Minimum step length=0.05
  Transform scale=100000.0
  Relaxation factor=0.9
  Maximum step length=10.0
  Initialize: useMomentsAlign
Results
  Success rate on sample images: 100% (no visible misalignment after registration), fully automatic
  Computation time: under 5 seconds

## Slide 4: CVH phantom data for Varian: CT to planning MR
Generic registration (BRAINS)
  Rigid (6 DOF)
  Number of samples=100000
  Minimum step length=0.05
  Transform scale=100000.0
  Relaxation factor=0.9
  Maximum step length=10.0
  Initialize: useMomentsAlign
Results
  Registration is successful, fully automatic
  Computation time: under 7 seconds

## Slide 5: CVH phantom data for Varian: CBCT to planning MR
Generic registration (BRAINS)
  Rigid (6 DOF)
  Number of samples=100000
  Minimum step length=0.05
  Transform scale=100000.0
  Relaxation factor=0.9
  Maximum step length=10.0
  Initialize: useMomentsAlign
Results
  Registration is successful, fully automatic
  Computation time: under 5 seconds

## Slide 6: Odette Cancer Centre data: CT to CBCT
Generic registration (BRAINS)
  Rigid (6 DOF)
  Number of samples=100000
  Minimum step length=0.05
  Transform scale=1000000.0
  Relaxation factor=0.9
  Maximum step length=5.0
  Initialize: None
Results
  Registration is successful, fully automatic
  Computation time: under 5 seconds